## Slide 1: Predicting Goldmann Applanation IOP
AIML PROJECT • BEC309L
Machine learning estimation of GAT IOP from iCare rebound tonometry and central corneal thickness (CCT)
6,639
eyes after cleaning
2
input features
5
models compared
iCare + CCT → ML → GAT
REGRESSION • OUTPUT IN mmHg
TEAM
Abhay Anand • Surya Vadhani • Saisrikant Hari (23BEC0446)
B.Tech Electronics and Communication Engineering • Vellore Institute of Technology
FACULTY
Dr. Suraj Prakash Sahoo
github.com/abhay04now/Icare-ML
Notes: Project goal: estimate the Goldmann (GAT) IOP an eye would have had, from its iCare reading and CCT. 6,639 cleaned eyes, two inputs, five models.

## Slide 2: The Problem
CLINICAL MOTIVATION
IOP is the main modifiable risk factor in glaucoma.
GAT is the clinical reference, but it needs a slit lamp, topical anaesthetic and a trained operator.
iCare is portable and needs no anaesthetic, but its readings differ systematically from GAT.
Central corneal thickness (CCT) is thought to explain part of that difference: thick corneas read high, thin corneas read low.
CORE QUESTION
Can machine learning predict GAT IOP from an iCare reading and CCT?
Notes: GAT is accurate but impractical outside the clinic. iCare is convenient but disagrees with GAT. The project asks whether ML can close that gap.

## Slide 3: Proposed ML Solution
FORMULATION
iCare IOP
INPUT
CCT
INPUT
→
ML REGRESSOR
Learns the mapping from the two inputs to measured GAT IOP
→
PREDICTED GAT IOP
Continuous output in mmHg
SCOPE
Age and gender are not in the dataset, so the model uses only iCare IOP and CCT.
REFERENCE POINTS
Every model must beat two non-ML baselines:
GAT = iCare, and GAT = 3.39 + 0.85 × iCare
Notes: Two inputs only. Age was in the original plan but is not available in the data. Any model is judged against reading iCare as-is and a straight line.

## Slide 4: Dataset & Cleaning Pipeline
DATA
01
Pair
iCare before GAT, same day, within 3 hours
02
De-duplicate
Exact copies of a measurement removed
03
Range filter
CCT 480–580 µm; IOP 0–80 mmHg
04
One per eye
Earliest valid observation kept
05
Patient split
Both eyes of a patient in the same split
6,639
eyes in total
4,648
train • 2,536 patients
1,001
validation
990
test
Row: date | euid | icare_iop | cct | applanation_iop Inputs in blue, target in green. Split ~70 / 15 / 15 by patient, random_state = 42.
Notes: Five sequential cleaning stages turn the raw hospital export into one row per eye. The split is by patient so no patient's eyes appear in both training and test.

## Slide 5: Why Regression?
AIML FORMULATION
REGRESSION
“What is the predicted GAT IOP value?”
Continuous value • mmHg
CLASSIFICATION
“Which class does this patient belong to?”
Discrete class label
GAT IOP is a number in mmHg, so the task is regression and the error is reported in mmHg.
Notes: The target is a continuous pressure, so regression is the right framing.

## Slide 6: Models Evaluated
MODEL FAMILIES
01
Linear Regression
iCare + CCT, least squares
02
Random Forest
100 trees, max depth 4
03
ANN / MLP
PyTorch 2-32-16-1, Adam
04
XGBoost
Depth 3, 500 trees, learning rate 0.01, λ = 10
05
SVR
RBF kernel, C = 15.4, γ = 0.18, ε = 0.25
BASELINES (NO MACHINE LEARNING)
Identity
GAT = iCare (report the device reading as-is)
Straight line
GAT = 3.39 + 0.85 × iCare
Notes: Five model families with their final tuned settings, plus two non-ML baselines that set the bar.

## Slide 7: Experimental Methodology
TRAINING PIPELINE
01
Clean & pair
Five-stage pipeline, 6,639 eyes
02
Define X / y
X: iCare IOP, CCT
y: GAT IOP
03
Split by patient
70 / 15 / 15; no patient in two splits
04
Baselines first
Identity and straight line set the bar
05
Tune on validation
Sweeps, then random search with GroupKFold(5) on patient
06
Test once
Held-out test scored a single time, after all choices
Notes: GroupKFold keeps both eyes of a patient in the same fold; plain KFold would leak one eye into training and flatter the score. The test set was never used to choose anything.

## Slide 8: Evaluation Metrics
MEASURING PERFORMANCE
MAE
Mean Absolute Error
MAE = (1/n) Σ |y − ŷ|
Headline metric, in mmHg
LOWER
RMSE
Root Mean Squared Error
RMSE = √( (1/n) Σ (y − ŷ)² )
Penalises large errors
LOWER
R²
Coefficient of Determination
R² = 1 − Σ(y − ŷ)² / Σ(y − ȳ)²
Share of variation explained
HIGHER
Bland–Altman
Agreement with GAT
limits = bias ± 1.96 · SD
Bias and 95% limits of agreement
NARROWER
Also reported: error by clinical IOP band, residual plots and permutation importance.
Notes: y is the measured GAT IOP, y-hat the prediction, y-bar the mean. MAE is the headline because it reads directly in mmHg.

## Slide 9: Model Comparison on the Test Set
RESULTS
Chart: Test MAE (mmHg, lower is better)
| Identity (GAT = iCare) | Linear (iCare only) | Linear (iCare + CCT) | Random Forest (depth 4) | SVR (RBF, tuned) | XGBoost (tuned)
Test MAE | 3.214 | 3.005 | 3.002 | 2.92 | 2.912 | 2.901
Model | Val MAE | Test MAE | Test RMSE | Test R²
XGBoost (tuned) | 2.809 | 2.901 | 3.986 | 0.730
SVR (RBF, tuned) | 2.829 | 2.912 | 4.084 | 0.717
Random Forest (depth 4) | 2.806 | 2.920 | 4.000 | 0.728
Linear (iCare + CCT) | 2.892 | 3.002 | 4.113 | 0.713
Linear (iCare only) | 2.883 | 3.005 | 4.114 | 0.712
Identity (GAT = iCare) | 2.996 | 3.214 | 4.352 | 0.678
ANN 2-32-16-1 | 2.847 | not run | not run | not run
XGBoost is best on every test metric, but it beats a straight line by only 0.10 mmHg. The ANN was scored on validation only.
Notes: Bars start at zero on purpose: the models are almost the same. XGBoost 2.901 vs straight line 3.005 mmHg. Source: results/metrics/model_comparison.csv.

## Slide 10: Key Findings
RESULTS
0.04 mmHg
Spread between four model families on validation
The limit is the two inputs, not the algorithm.
0.10 mmHg
Gain of the best model over a straight line on test
Inside the 1–2 mmHg variability of GAT itself.
+0.007 mmHg
MAE increase when CCT is shuffled (iCare: +3.42)
CCT adds nothing within the 480–580 µm range.
> 30 mmHg
Every model is worse than the raw iCare reading
Models under-predict high pressure, where errors matter most.
Notes: Four model families converge; CCT importance near zero by permutation, ablation and linear fit; high-pressure eyes are the failure case.

## Slide 11: Error by Clinical IOP Band
RESULTS
Chart: Validation MAE (mmHg) by GAT IOP band (eyes)
| <10 (15) | 10–20 (641) | 20–30 (267) | 30–40 (55) | 40+ (23)
XGBoost | 3.73 | 2.09 | 3.11 | 6.76 | 9.24
SVR | 3.26 | 1.97 | 3.4 | 7.25 | 9.26
Raw iCare | 1 | 2.46 | 3.3 | 5.91 | 8.78
WHAT IT SHOWS
SVR is best in the 10–20 mmHg band (1.97 vs 2.46 raw iCare).
Above 30 mmHg both models are worse than raw iCare and under-predict.
Cause: 93% of training eyes are below 30 mmHg.
Notes: Overall MAE is dominated by the 10-20 band. Breaking it out shows the failure at high pressure, which is clinically the important range.

## Slide 12: Agreement with GAT on the Test Set
RESULTS
XGBoost
Bias
+0.37 mmHg
95% limits
−7.41 to +8.16 mmHg
Within ±2 mmHg
47.0%
Within ±5 mmHg
86.0%
SVR
Bias
+0.83 mmHg
95% limits
−7.02 to +8.67 mmHg
Within ±2 mmHg
48.6%
Within ±5 mmHg
82.9%
ROUNDING TO THE GAT GRID
78% of GAT readings are even numbers: the dial is read in 2 mmHg steps.
47% → 60%
within ±2 mmHg when a quadratic linear model's predictions are rounded to the nearest even number (MAE 2.909 → 2.867).
Notes: Limits of agreement are about ±8 mmHg for both models. The rounding result comes from notebooks/icare_experiments.ipynb.

## Slide 13: Conclusion & Next Steps
CONCLUSION
WHAT WE ESTABLISHED
Best model: XGBoost, test MAE 2.901 mmHg, R² 0.730.
ML beats raw iCare by 0.31 mmHg and a straight line by 0.10 mmHg.
With only iCare and CCT, the information is exhausted: five methods agree.
Not ready for clinical use: high-pressure eyes are under-predicted.
NEXT STEPS
CCT RANGE
Widen the 480–580 µm filter and repeat the CCT ablation.
MORE DATA
Keep repeat visits per eye instead of dropping ~36% of rows.
HIGH IOP
Fix the bias with sample weighting or monotonic constraints.
NEW INPUTS
Add age, corneal hysteresis and axial length.
ANN
Score the neural network on the test set.
Notes: Main message: the algorithm is not the bottleneck, the inputs are. The next gains come from data, not from more tuning.